## Slide 1
H2020-MSCA-RISE-2015 ⎯ Grant Agreement N° 690835
Muse Collaboration Meeting,
October 23, 2018
Analysis of Source and Local Monitor Signals.
Nandita Raha, INFN Pisa

## Slide 2: Laser System - Basics
Laser system - used for calibration and monitoring.
  Source Monitors (SM) are used to monitor and correct the laser intensity variations. Have 2 PIN diodes (PiD) and a PMT.
      Simultaneous digitization with two independent systems: Waveform digitizers (WFDs) similar to those used for the calorimeters and custom (Naples) electronics.
  Local Monitors (LM) check and correct the stability of the light distribution chain. The same light pulse is split into two separate paths: one is measured locally from the SM, while the other goes to the calorimeter and then returns to the laser hut. The two signals are measured by the same PMT and clearly separated in time.

## Slide 3
Laser System - Basics
SM: Monitors the laser source.
x 6

## Slide 4: Laser Monitoring System
Map of the laser monitoring system from reconstruction. The red block shows the 17 SMs (SM6 PMT not working shown by red cross and channel 15, 20 and 19 are not connected shown by blue crosses) and the rest show the 24 LMs.
LM
SM

## Slide 5: Parameter definitions
Pedestal: Average first 50 samples
Amplitude: Pedestal subtracted peak sample
Area window: + 50 samples from peak.
Note: Same definitions for Source Monitors too.
Sample number

## Slide 6: Offline Studies
Studies of SMs with WFDs and comparison with custom (Naples) electronics:
  Signal amplitude
  Stability of SMs.
LM studies:
  Signal amplitude
  Stability of light chain (amplitude ratio of the 2 LM pulses)
  PMT Pedestal
  Time interval between the two LM pulses

## Slide 7: Time Evolution Studies Runs 1791 to 1869 (1st to 4th July) Source Monitors

## Slide 8: SM from WFDs
Scatter plot between amplitudes of PiD1 and PiD2 shows a strong correlation, for SM4 which, follow the laser light fluctuations.

## Slide 9: Temperature Dependence
Laser hut room temperature – read by Slow control
Custom Naples Board of the PIN 1
The amplitudes depends both on the ambient temperature (of the laser hut) and on the PiD read-out boards. The board temperature(43 oC) is higher than the room temperature, but its variation follows the trend of the room temperature.
Gain variation -0.4%/ oC

## Slide 10: Source Monitor 4
WFDs
Custom Naples Electronics
The ratio of amplitudes of PiD1 to PiD2 is a measure of the relative stability of the front end and readout electronics. Compared with two independent systems i.e. WFD’s and custom Naples electronics. Lines correspond to an ~10-4.

## Slide 11: Source Monitor 5
The relative stability is not very good here. Need to investigate the reason for the drifts here. We can’t compare with the custom Naples electronics as SM1, SM5 and SM6 could not be read by their boards.

## Slide 12: Summary of SM
Aim to achieve a stability of ~ 10-4 (stability is measured by the amplitude ratio of PiD1:PiD2).
Check this stability with two independent systems i.e. WFDs and custom Naples electronics .
We see local drift for certain PiDs which we are investigating.

## Slide 13: Time Evolution Studies Runs 1791 to 1869 (1st to 4th July) Local Monitors

## Slide 14: LM Amplitudes
Scatter plot between amplitude 1 and amplitude 2 – almost no correlation for a particular channel – selected this for investigation. PMT fluctuations are larger than laser fluctuation.

## Slide 15
Stability checks – One channel
Time difference: t2 - t1
A2/A1
Pedestal

## Slide 16
Stability Checks for a Channel
Pedestal
Perform this on all channels and study the variation of sigma (spread) to check stability of pedestal and amplitudes of both pulses. We also investigated the time difference t2-t1 for all channel (did not show it since we have just three time bins).
Sigma 0.72 counts
A2/A1 wrt mean
Sigma 5.7%

## Slide 17
Area Studies
Fitted area distribution of all waveforms of each channel with a Gaussian and plotted the mean values and the reduced c2 . 8th local monitor channel and second pulse area (right plot). Same for Frascati test beam (left). A broader spread un this run is due to smaller number of photo electron arriving at the calorimeter. The main effect corresponds to a factor of on the s (as observed).
s/mean= 2.2%
Courtesy: A. Druitti
Equivalent pulse energy ~ 2 GeV
Equivalent pulse energy ~ 1 GeV
s/mean= 2.2%
s/mean= 3.4%
Runs 1791 to 1869

## Slide 18
Area A2 Width Vs. channels
Shows the sigma/mean (width) for all channel.

## Slide 19
Pedestal Width Vs. channels
Shows the sigma (width) for all channel.

## Slide 20: Summary of LM
  Ratio s/mean of area of second pulse varies from 3.4 to 5.9%, higher than the Frascati test beam (~2.5-3%) . A beam of laser light equivalent of 1 GeV compared to Frascati beam (of ~2 GeV), introduces a reduction by a factor of in the width (i.e. ratio s/mean) which depends on the number of photoelectrons (PE) arriving on the calorimeters .

## Slide 21: Conclusion
  The stability of the SMs measured by amplitude ratio of PiD1:PiD2 is generally good (of the ~10-4), but some noise and temperature dependence have been observed, which we need to understand.
  LM don’t see too much fluctuations. Optimization of light will stabilize and improve their behaviour.
  Redefining the parameters of the pulse ( like pedestals, amplitudes, areas etc.) using a template fit or any other sophisticated fits might help in achieving a better stability of both the SMs and LMs.

## Slide 22: Backup Slides

## Slide 23
The amplitude and area are proportional as shown 1st pulse (left) and 2nd pulse(right) for channel 1.
LM: Amplitude Vs. Area
Runs 1791 to 1869
1st – 4th July

## Slide 24
Amplitude Studies – First pulse
Error is 1/[2*sqrt(NDF)]. Selected channel 10 (calorimeter 15) with least reduced c2 for time evolution studies.
Within a fill
The Gaussian distribution of amplitudes for each channel is fitted and the reduced c2 distribution is plotted to chose a suitable channel.

## Slide 25: SM1 from WFD’s
The plot of ratio of pin diodes amplitudes only for source monitor 1 is shown here.
All other monitors have similar behaviour

## Slide 26: SM2 from WFD’s