## Slide 1
Nandita Raha, INFN Pisa
International School on Muon Dipole Moments and Hadronic Effects
Calibration System of the the Muon g-2 Experiment at Fermilab

## Slide 2: Outline
Theoretical Background and Basic Technique
Experimental Setup and Systematics
Laser Calibration System
Local Monitors of the Laser Calibration System
Stability of the Local Monitors – time evolution studies
Stability of the Local Monitors – temperature studies
Conclusions
Notes: 2

## Slide 3: Muon g-2: The Basics
3.7 s
7 s
gSM + gNP? =
+
+
+
+
NP
3
with g = 2 for a tree level anomaly (Dirac)
g = 2(Dirac) + O(10-3)QED + O(10-9)EW + O(10-7)QCD + NP (?)
Notes: 3

## Slide 4: Muon g-2: The Experiment
Magic p = 3.09 GeV at g = 29.3
B
4
with g = 2 for a tree level anomaly (Dirac)
Notes: 4

## Slide 5: Muon anomalous moment measurement
E821 data: e+ with E > 1.8 GeV
Systematic effects in the positron spectrum due to:
  Pileup, gain stability
  Beam Effects, Muon losses
5
wa
wp
Measure B  via NMR  recast am in terms of proton precession frequency, wp (at B in its rest frame).
Muonic hyperfine experiment gives mm/mp at ~26 ppb precision (ref. arXiv:1203.5425)

## Slide 6
Systematic Improvements on wa
Systematics on ωa 180 70 ppb compared to E821. Achieved by:
  Improved Calorimeters
  Much improved Laser control system
  New Straw Tracker (two installed in run1)
Key element:
Laser
Calo + Laser
Calo + Laser
Inflector + Kicker
Tracker

## Slide 7: Monitoring / Calibration of Calorimeters - Laser
7
Laser calibration system –Optical fibers carry laser shots to each of the 1296 calorimeter crystals. One laser head sends light to 4 calos as shown in the right fig.
6 SM (source monitors - 2 pin diodes + 1 PMT) to check and correct for laser light fluctuations.
24 LM (local monitors) each with two PMTs that are used to check stability of the light distribution chain.
The picture on left shows the entire laser calibration system in the laser hut of the experimental hall.
Notes: Make more slides – three points – explaining – Laser light on calo for calibrating them. SM to check fluctuations of Laser light and LM to check stability of the optical system

## Slide 8
WFD g-2
Local DAQ & eventually to WFD g-2
PMMA Fibre
Quartz Fibre
Foward
Backward
A2=L*F*B, F= forward eff; B=Back eff
If F Gain changes SiPM should be corrected; IF B no
correction depends on R=A2/A1
A2
Local Monitors - details

## Slide 9: LM Analysis – time evolution studies for 60 hours dataset 22/4 - 1:09 pm to 25/4 – 2:20 am

## Slide 10: Comparison of old and new PMTs for Calo 20
New PMT – Calo 20, channel 4
Old PMT – Calo 20, channel 29
Bin size = 50 mins.

## Slide 11: Temperature Variation – Channel 4
Overlaid temperature and A1, A2 and A2:A1 on the same plot, after normalizing all to A1 for consistency
Optical table temperature
Temperature (oC)

## Slide 12: A1 and A2 overlaid and A2:A1 – 20 hr test dataset
PMMA Channel 23
Silica Channel 24
A2:A1
A2:A1

## Slide 13: LM Analysis – temperature studies of 60 hr dataset 22/4 - 1:09 pm to 25/4 – 2:20 am

## Slide 14: Amplitude Vs. Temp – Channel 4
The slope appears to change at 29.8 C.
Temperature (oC)
Temperature (oC)
Temperature (oC)
A1
A2
A2:A1

## Slide 15: Amplitude Vs. Temp – channels 1 - 24
Amp 1
Amp 2
A2:A1
Results obtained by linear fits of each channel.

## Slide 16: 100 hours dataset – Temperature Vs. A1 – Channel 4
A1 (Normalized)
Temperature (oC)
A1 normalized
A different dataset confirming the negative temperature coefficient of gain with temperature. Around the temperature of 32.65 oC, we had a couple of bad runs which causes a small distortion (red circle below).
Thus, a cross check with another dataset and another channel confirms similar results

## Slide 17
Temperature studies of 100 hour dataset
New PMT – Silica fiber both forward and return directions
A1
A2
A2:A1

## Slide 18: Conclusions
Variation of A1 with temperature shows a negative correlation of PMT gain with temperature (as expected).
A2:A1 is almost constant with time for the (for longer time ranges some variation appears especially for the old PMTs)
Channels 0 to 23 i.e. the new PMTs (refer slide 9) show a comparatively stable A2:A1 drop with temperature (it ranges from ~0.1% to 0.45%) which is much better than the old PMTs and confirms a reasonably stable behaviour (in gain) which meets the required precision of the experiment.
Thank You!!

## Slide 19: Back up slides

## Slide 20: Back up and other basic info
About laser used – PicoQuant(LDH-P-C-405M), max pulse energy 1 nJ, pulse width 700 ps @ wavelength 405+10 nm max. repetition rate 40 MHz.
Beam fill structure: Fill is 600 us each. Has 3 out of fill windows of 10 ms, ~200 ms and ~1000 ms
Pulse quality display in next slides

## Slide 21: Calorimeters
24 calo – each divided in a grid of 9x6 PbFl Cerenkov crystals read out with SiPM – 1 calo 54 SiPM (for better pileup) @ 800 MSPS, uTCA crates read by a MIDAS/U DAQ
energy res of 50 MeV at 2 GeV.
Pile up separation 4.5 ns and time res 25 ps at 3 GeV
Additionally have 3 straw trackers 128 straws Ar:CO2 gas – would help track positrons

## Slide 22
Some Studies/Observations –Scatter plot and projection of a bin – good Gaussian

## Slide 23
Correlation between Ped / amplitude ratio

## Slide 24: Revisit temperature
A1 with temp overlaid. Again shows negative correlation. Analyzed till 1100 mins. Explains plots of slide 12.